# Calculation of effective dose due to exposure to a point source

- Discussion and equations

Effective doses calculated at 1 m from the point source are used to calculate the dose at other distances, assuming no shielding.  For distances shorter than 0.5 m, the calculation is will significantly overestimate the real dose rate. For distances less than 0.2 m, the equation should not be used. The equation is:

Epoint=A x CFpoint x tdistance2

Where

Epoint	=	Effective dose from point source [mSv]

A =	Source activity [kBq]. 1 Ci = 3.7 x 107 kBq.

CFpoint	= Conversion factor for a point source, see section 2 [(mSv . m2)/(kBq . h)]

distance =	Distance in meters from the point source [m]

t =	Exposure duration [h]

When information on the source shielding (material and thickness) is available, the calculated effective dose from the point source may be reduced by using the half value layer  (HVL) for the material in question. The half value layer is the thickness required to reduce the dose by a half. The following equation should be used:

Epointshielded= Epoint x (0.5)thicknessHVL

Where

Epoint shielded	=	Effective dose from shielded point source [mSv]

thickness = material thickness [cm]

HVL = half value layer for the material [cm], see the HVL table in section 3.

2) Conversion factors

| Radionuclide | CFpoint (mSv . m2)/(kBq . h) |
| --- | --- |
| Na-22 Na-24 | 2.2E-07 3.8E-07 |
| K-40 K-42 | 1.6E-08 2.8E-08 |
| Sc-46 | 2.1E-07 |
| Ti-44 | 1.1E-08 |
| V-48 | 2.9E-07 |
| Cr-51 | 3.4E-09 |
| Mn-54 Mn-56 | 8.6E-08 1.7E-07 |
| Fe-55 Fe-59 | 3.2E-10 1.2E-07 |
| Co-58 Co-60 | 1.0E-07 2.5E-07 |
| Cu-64 | 2.0E-08 |
| Zn-65 | 6.0E-08 |
| Ga-68 | 9.8E-08 |
| Ge-68+Ga-68 | 9.8E-08 |
| Se-75 | 3.9E-08 |
| Kr-85 Kr-85m Kr-87 Kr-88+Rb-88 | 2.3E-10 1.5E-08 7.8E-08 2.5E-07 |
| Rb-86 Rb-88 | 9.6E-09 5.7E-08 |
| Sr-89 Sr-91 | 1.4E-11 7.1E-08 |
| Y-91 Y-91m | 3.7E-10 5.5E-08 |
| Zr-95 | 7.6E-08 |
| Nb-94 Nb-95 | 1.6E-07 7.9E-08 |
| Mo-99 | 1.6E-08 |
| Tc-99m | 1.2E-08 |
| Rh-103 | 2.1E-08 |
| Ru-103 Ru-105 Ru-106 | 5.0E-08 8.1E-08 1.4E-09 |
| Ru-106+Rh-106 | 1.4E-09 |
| Ag-110m | 2.8E-07 |
| Cd-109+Ag-109m | 1.6E-07 |
| In-114m | 1.0E-08 |
| Sn-113 Sn-123 | 3.4E-09 7.0E-10 |
| Sn-126+Sb-126m | 5.7E-09 |
| Sb-124 Sb-126 Sb-126m Sb-127 Sb-129 | 1.9E-07 2.8E-07 4.9E-10 6.8E-08 1.5E-07 |
| Te-127 Te-127m Te-129 Te-129m+Te-129 Te-131 Te-131m Te-132 | 6.0E-09 1.6E-09 4.2E-08 4.6E-08 4.5E-08 1.5E-07 2.3E-08 |
| I-125 I-129 I-131 I-132 I-133 I-134 | 5.9E-09 3.4E-09 3.9E-08 2.4E-07 6.2E-08 2.7E-07 |
| I-135+Xe-135 | 3.8E-07 |
| Xe-131m Xe-133 Xe-133m Xe-135 Xe-138 | 2.7E-09 4.6E-09 4.8E-09 2.4E-08 1.1E-07 |
| Cs-134 Cs-136 Ba-137m Cs-137+Ba-137m Ba-133 Cs-138 Ba-140 | 1.6E-07 2.2E-07 6.2E-08 6.2E-08 4.1E-08 3.0E-09 2.0E-08 |
| La-140 | 2.3E-07 |
| Ce-141 Ce-144+ Pr-144 | 7.2E-09 3.1E-09 |
| Pr-144m Pr-144 | 2.9E-09 1.2E-09 |
| Pm-145 | 3.6E-09 |
| Eu-152 Eu-154 Eu-155 | 1.2E-07 1.3E-07 5.3E-09 |
| Gd-153 | 1.1E-08 |
| Tb-160 | 1.1E-07 |
| Ho-166m | 1.6E-07 |
| Tm-170 | 5.0E-10 |
| Yb-169 | 2.9E-08 |
| Hf-172 Hf-181 | 2.2E-08 5.5E-08 |
| Ta-182 | 1.3E-07 |
| W-187 | 4.9E-08 |
| Ir-192 | 8.3E-08 |
| Au-198 | 4.1E-08 |
| Hg-203 | 2.3E-08 |
| Ti-204 | 1.0E-10 |
| Pb-210 | 6.9E-10 |
| Bi-207 | 1.6E-07 |
| Ra-226 | 6.2E-10 |
| Ac-228 | 9.5E-08 |
| Th-227 Th-228 Th-230 Th-231 Th-232 | 1.1E-08 3.9E-10 2.3E-10 2.5E-10 2.1E-10 |
| Pa-231 | 4.3E-09 |
| U-Dep & Nat U-Enriched U-232 Pa-233 U-233 U-234 U-235 U-238 | 2.3E-10 2.8E-10 3.2E-10 1.7E-08 1.2E-10 2.8E-10 1.4E-08 2.3E-10 |
| Np-237 Pu-236 Pu-238 Pu-239 Pu-240 Pu-242 | 3.8E-09 3.4E-10 3.0E-10 1.2E-10 2.8E-10 2.3E-10 |
| Am-241 Am-242 Am-243 | 3.1E-09 8.5E-10 5.4E-09 |
| Cm-242 Cm-243 Cm-244 Cm-245 | 3.1E-10 1.3E-08 2.8E-10 7.5E-09 |
| Cf-252 | 2.1E-10 |

Calculations made by Oak Ridge National Laboratory (ORNL) using the CONDOS programme. All other radionuclides have a CFpoint  < 10 -10 (mSv . m2)/(kBq . h) and need not be considered.  Information based on IAEA TECDOC – 1162: Generic procedures for assessment and response during a radiological emergency. IAEA, August 2000.

3) Half Value Layers

The half value layer (HLV) is the thickness of a substance that, when introduced in the path of a beam of radiation, reduces the exposure rate by one-half. Values are given for "good geometry" for which build-up of secondary radiation is not important.

| Radionuclide | Half Value Layer (cm) |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Lead | Iron | Al | Water | Concrete |
| Na-22 | 0.67 | 1.38 | 3.85 | 9.4 | 4.35 |
| Na-24 | 1.32 | 2.14 | 6.22 | 14.75 | 6.88 |
| K-40 | 1.15 | 1.8 | 4.99 | 11.97 | 5.63 |
| K-42 | 1.18 | 1.84 | 5.1 | 12.21 | 5.75 |
| Sc-46 | 0.82 | 1.48 | 4.2 | 9.84 | 4.66 |
| Ti-44 | 0.04 | 0.21 | 0.6 | 1.41 | 0.67 |
| V-48 | 0.8 | 1.48 | 4.18 | 9.95 | 4.67 |
| Cr-51 | 0.17 | 0.82 | 2.38 | 5.69 | 2.68 |
| Mn-54 | 0.68 | 1.33 | 3.8 | 9.0 | 4.22 |
| Mn-56 | 0.94 | 1.65 | 4.78 | 11.13 | 5.27 |
| Fe-59 | 0.94 | 1.59 | 4.51 | 10.58 | 5.02 |
| Co-60 | 1 | 1.66 | 4.65 | 10.99 | 5.2 |
| Cu-64 | 0.41 | 1.08 | 3.01 | 7.61 | 3.43 |
| Zn-65 | 0.87 | 1.53 | 4.34 | 10.15 | 4.81 |
| Ga-68 | 0.42 | 1.09 | 3.04 | 7.67 | 3.47 |
| Ge-68+Ga-68 | 0.42 | 1.09 | 3.04 | 7.67 | 3.47 |
| Se-75 | 0.12 | 0.62 | 1.79 | 4.26 | 2.01 |
| Kr-85 | 0.41 | 1.07 | 3 | 7.59 | 3.43 |
| Kr-85m | 0.1 | 0.5 | 1.46 | 3.46 | 1.64 |
| Kr-87 | 0.83 | 1.67 | 4.84 | 11.46 | 5.36 |
| Kr-88+Rb-88 | 1.17 | 1.89 | 5.51 | 12.74 | 6.05 |
| Rb-86 | 0.87 | 1.53 | 4.35 | 10.13 | 4.81 |
| Rb-88 | 1.17 | 1.89 | 5.51 | 12.74 | 6.05 |
| Sr-89 | 0.74 | 1.4 | 4 | 9.35 | 4.42 |
| Sr-91 | 0.71 | 1.38 | 3.94 | 9.31 | 4.38 |
| Y-91 | 0.96 | 1.62 | 4.57 | 10.74 | 5.09 |
| Zr-95 | 0.6 | 1.26 | 3.58 | 8.61 | 4.0 |
| Nb-94 | 0.64 | 1.30 | 3.70 | 8.84 | 4.13 |
| Nb-95 | 0.62 | 1.28 | 3.63 | 8.72 | 4.06 |
| Mo-99+Tc-99m | 0.49 | 1.11 | 3.16 | 7.6 | 3.54 |
| Mo-99 | 0.49 | 1.11 | 3.16 | 7.6 | 3.54 |
| Tc-99 | 0.05 | 0.25 | 0.73 | 1.73 | 0.82 |
| Tc-99m | 0.07 | 0.39 | 1.13 | 2.68 | 1.27 |
| Ru-103 | 0.4 | 1.06 | 2.97 | 7.53 | 3.4 |
| Ru-105 | 0.48 | 1.16 | 3.28 | 7.98 | 3.69 |
| Rh-106 | 0.49 | 1.17 | 3.29 | 8.16 | 3.73 |
| Ag-110m | 0.71 | 1.38 | 3.91 | 9.36 | 4.38 |
| Cd-109 | 0.01 | 0.06 | 0.18 | 0.43 | 0.2 |
| In-114m | 0.23 | 0.75 | 2.14 | 5.18 | 2.41 |
| Sn-113 | 0.02 | 0.09 | 0.27 | 0.65 | 0.31 |
| Sn-123 | 0.88 | 1.53 | 4.36 | 10.16 | 4.83 |
| Sn-126+Sb-126m | 0.48 | 1.15 | 3.27 | 7.99 | 3.68 |
| Sn-126 | 0.04 | 0.19 | 0.55 | 1.3 | 0.62 |
| Sb-124 | 0.83 | 1.55 | 4.39 | 10.49 | 4.9 |
| Sb-126 | 0.52 | 1.19 | 3.37 | 8.21 | 3.79 |
| Sb-126m | 0.48 | 1.15 | 3.27 | 7.99 | 3.68 |
| Sb-127 | 0.47 | 1.14 | 3.24 | 7.92 | 3.65 |
| Sb-129 | 0.72 | 1.4 | 3.98 | 9.45 | 4.43 |
| Te-127m | 0.01 | 0.08 | 0.23 | 0.54 | 0.26 |
| Te-129 | 0.33 | 0.93 | 2.63 | 6.53 | 2.99 |
| Te-129m | 0.38 | 0.82 | 2.33 | 5.65 | 2.61 |
| Te-131m | 0.65 | 1.31 | 3.74 | 8.88 | 4.17 |
| Te-132 | 0.1 | 0.53 | 1.54 | 3.66 | 1.73 |
| I-125 | 0.01 | 0.08 | 0.23 | 0.54 | 0.26 |
| I-129 | 0.02 | 0.09 | 0.25 | 0.6 | 0.28 |
| I-131 | 0.25 | 0.93 | 2.67 | 6.5 | 3.02 |
| I-132 | 0.63 | 1.31 | 3.7 | 8.91 | 4.14 |
| I-133 | 0.47 | 1.15 | 3.23 | 8.05 | 3.67 |
| I-134 | 0.72 | 1.4 | 3.98 | 9.43 | 4.43 |
| I-135+Xe-135m | 0.98 | 1.66 | 4.7 | 11.06 | 5.23 |
| I-135 | 0.98 | 1.66 | 4.7 | 11.06 | 5.23 |
| Xe-131m | 0.02 | 0.1 | 0.29 | 0.7 | 0.33 |
| Xe-133 | 0.03 | 0.16 | 0.47 | 1.11 | 0.53 |
| Xe-133m | 0.05 | 0.25 | 0.73 | 1.72 | 0.82 |
| Xe-135 | 0.14 | 0.72 | 2.1 | 4.99 | 2.36 |
| Xe-135m | 0.41 | 1.07 | 2.99 | 7.54 | 3.41 |
| Xe-138 | 0.9 | 1.64 | 4.79 | 11.09 | 5.26 |
| Cs-134 | 0.57 | 1.24 | 3.5 | 8.5 | 3.93 |
| Cs-136 | 0.65 | 1.32 | 3.76 | 8.86 | 4.18 |
| Cs-137+Ba-137m | 0.53 | 1.19 | 3.35 | 8.2 | 3.77 |
| Ba-133 | 0.16 | 0.67 | 1.92 | 4.63 | 2.17 |
| Ba-137m | 0.53 | 1.19 | 3.35 | 8.2 | 3.77 |
| Ba-140 | 0.33 | 0.96 | 2.69 | 6.72 | 3.06 |
| La-140 | 0.93 | 1.64 | 4.63 | 11.04 | 5.19 |
| Ce-141 | 0.07 | 0.37 | 1.07 | 2.52 | 1.2 |
| Ce-144+Pr-144m | 0.05 | 0.28 | 0.82 | 1.95 | 0.93 |
| Pr-144m | 0.02 | 0.1 | 0.28 | 0.67 | 0.32 |
| Pm-145 | 0.02 | 0.11 | 0.31 | 0.74 | 0.35 |
| Pm-147 | 0.06 | 0.34 | 0.99 | 2.35 | 1.12 |
| Sm-151 | 0.01 | 0.03 | 0.09 | 0.21 | 0.1 |
| Eu-152 | 0.66 | 1.32 | 3.73 | 8.84 | 4.17 |
| Eu-154 | 0.74 | 1.38 | 3.91 | 9.24 | 4.35 |
| Eu-155 | 0.04 | 0.23 | 0.66 | 1.56 | 0.74 |
| Gd-153 | 0.03 | 0.18 | 0.51 | 1.21 | 0.57 |
| Tb-160 | 0.68 | 1.35 | 3.84 | 9.01 | 4.26 |
| Ho-166m | 0.45 | 1.09 | 3.1 | 7.46 | 3.48 |
| Tm-170 | 0.03 | 0.18 | 0.51 | 1.21 | 0.57 |
| Yb-169 | 0.06 | 0.3 | 0.87 | 2.05 | 0.97 |
| Hf-181 | 0.27 | 0.86 | 2.41 | 6.02 | 2.75 |
| Ta-182 | 0.8 | 1.39 | 3.94 | 9.26 | 4.39 |
| W-187 | 0.43 | 1.03 | 2.91 | 7.17 | 3.29 |
| Ir-192 | 0.24 | 0.92 | 2.64 | 6.42 | 2.98 |
| Au-198 | 0.29 | 0.97 | 2.74 | 6.77 | 3.11 |
| Hg-203 | 0.14 | 0.73 | 2.13 | 5.04 | 2.39 |
| Tl-204 | 0.03 | 0.18 | 0.53 | 1.27 | 0.6 |
| Pb-210 | 0.01 | 0.05 | 0.15 | 0.35 | 0.17 |
| Bi-207 | 0.65 | 1.3 | 3.68 | 8.79 | 4.11 |
| Po-210 | 0.65 | 1.31 | 3.73 | 8.88 | 4.15 |
| Ra-226 | 0.09 | 0.48 | 1.4 | 3.32 | 1.58 |
| Ac-227 | 0.01 | 0.08 | 0.22 | 0.52 | 0.25 |
| Ac-228 | 0.67 | 1.35 | 3.84 | 9.05 | 4.27 |
| Th-227 | 0.11 | 0.58 | 1.69 | 4.01 | 1.9 |
| Th-228 | 0.02 | 0.13 | 0.37 | 0.88 | 0.42 |
| Th-230 | 0.01 | 0.05 | 0.14 | 0.34 | 0.16 |
| Th-232 | 0.01 | 0.04 | 0.12 | 0.28 | 0.13 |
| Pa-231 | 0.09 | 0.46 | 1.35 | 3.2 | 1.51 |
| U-232 | 0.01 | 0.04 | 0.12 | 0.29 | 0.14 |
| U-233 | 0.01 | 0.06 | 0.16 | 0.39 | 0.18 |
| U-234 | 0.01 | 0.04 | 0.12 | 0.28 | 0.13 |
| U-235 | 0.09 | 0.46 | 1.35 | 3.19 | 1.51 |
| U-238 | 0.01 | 0.04 | 0.11 | 0.27 | 0.13 |
| Np-237 | 0.03 | 0.12 | 0.41 | 0.98 | 0.46 |
| Pu-236 | 0.01 | 0.04 | 0.11 | 0.27 | 0.13 |
| Pu-238 | 0.01 | 0.04 | 0.11 | 0.27 | 0.13 |
| Pu-239 | 0.01 | 0.04 | 0.12 | 0.29 | 0.14 |
| Pu-240 | 0.01 | 0.04 | 0.11 | 0.27 | 0.13 |
| Pu-242 | 0.01 | 0.04 | 0.11 | 0.27 | 0.13 |
| Am-241 | 0.02 | 0.12 | 0.35 | 0.82 | 0.39 |
| Am-242m | 0.01 | 0.04 | 0.13 | 0.3 | 0.14 |
| Am-243 | 0.03 | 0.18 | 0.52 | 1.24 | 0.59 |
| Cm-242 | 0.01 | 0.04 | 0.12 | 0.28 | 0.13 |
| Cm-243 | 0.08 | 0.43 | 1.26 | 2.98 | 1.41 |
| Cm-244 | 0.01 | 0.04 | 0.12 | 0.28 | 0.13 |
| Cm-245 | 0.05 | 0.27 | 0.79 | 1.86 | 0.88 |
| Cf-252 | 0.01 | 0.04 | 0.12 | 0.3 | 0.14 |

## 4) Example

A worker remained one hour at a distance of 5 meters from a 27 Ci 192Ir source. What was his effective dose?

Ir-192 : CFpoint = 8.3E-08 (mSv . m2)/(kBq . h)

Epoint=A x CFpoint x td2

Epoint=27 x 3.7 x 107 x 8.3 x 10−8 x 152

Epoint	= 3.3 mSv.

Total effective dose from this source at 5 meters in 1 hour = 3.3 mSv.

If the Ir-192 source was shielded with 5 mm of lead what is the dose calculated with the shielding taken into account?

Epointshielded= Epoint x (0.5)thicknessHVL

Ir-192: HVL for lead = 0.24 cm.

Epointshielded= 3.3 x (0.5)0.50.24

Epointshielded= 3.3 x 0.236 mSv

Epointshielded= 0.78 mSv

5) References

INTERNATIONAL ATOMIC ENERGY AGENCY, Generic Procedures for Assessment and Response during an Radiological Emergency, IAEA TECDOC- 1162, IAEA, Vienna (2000).