CONSIDERATIONS FOR RELEASE OF PATIENTS ADMINISTERED

RADIOPHARMACEUTICALS

Since 1997, members of the public have been limited to 5 mSv (500 mrem) Total Effective Dose Equivalent (TEDE) from exposure to appropriately released patients who have been administered radiopharmaceuticals, pursuant to 10 CFR 35.75. Many Agreement States have adopted the Nuclear Regulatory Commission’s approach to patient release.

The educational review material in this tutorial was compiled by Carol S. Marcus (csmarcus@ucla.edu), Jeffry A. Siegel (nukephysics@comcast.net) and Michael G. Stabin (michael.g.stabin@vanderbilt.edu). Comments, criticisms, and additions are welcome. Some of this material has already been posted on the internet at the request of the California Radiological Health Branch (RHB). These educational materials, posted by RADAR members, may be used by all Authorized Users (AUs), Radiation Safety Officers and other professionals involved in the administration of radionuclide therapy in nuclear medicine.

There are three major elements involved in successfully meeting a performance standard of maintaining exposure of members of the public to released nuclear medicine patients to under 5 mSv (500 mrem):

- The first is an evaluation of the patient’s living and working conditions to ascertain whether or not a given patient can be safely released.

- The second step is the appropriate performance of a patient-specific dose calculation to ensure that no individual member of the public will likely be exposed to a dose in excess of 5 mSv (500 mrem).

- The third is to provide verbal and written instructions that are simple in order for the patient to limit the radiation dose to others to as low as reasonably achievable (ALARA). This requires patient education and an assessment by the AU physician that patient compliance with these instructions is highly likely.

PATIENT EDUCATION

Physicians will need to tailor advice to the patient's educational level, cultural background, and socioeconomic circumstances. If for any reason the patient cannot be appropriately educated, or there is a reasonable expectation of patient non-compliance, the AU physician should consider hospitalizing the patient.

Control of body fluid contamination is of some concern with 131I-NaI and to a lesser extent with 131I-Bexxar, but is of minor concern with 89Sr-Cl2 (Metastron), 153Sm-EDTMP (Quadramet), 32P-chromic phosphate, 32P-sodium phosphate, and 90Y-Zevalin. External radiation exposure is an important consideration for the photon emitters 131I  and 153Sm.  Bremsstrahlung radiation associated with pure beta-emitting radionuclides is not a significant radiation protection concern as a source of external radiation absorbed dose to others.

Patients who receive 131I  and who will probably cook when they go home should be given several pairs of disposable gloves and taught how to wash and dry them, recalling that it is the inside, rather than the outside, which is of contamination concern. The usual caveats about kissing and sharing eating utensils and drinking vessels apply. The only new development here is the common installation of NaI radiation detectors in transfer stations and sanitary landfills which are able to pick up very small quantities of 131I and 153Sm, often undetectable with conventional nuclear medicine survey equipment. For this reason, patients should be told that whatever 131I and 153Sm cannot go out via the plumbing (toilet, sink, dishwasher, clothes washer) should be collected for a week and stored in a double plastic bag and returned to the physician administering the radiopharmaceutical.  Alternatively, the materials may be collected for a week or so and then discarded. Concerns about the radiation risks of such discarded materials are not generally proportional to the actual level of hazard.

Issues concerning return to work are generally important only with 131I patients, but can affect 153Sm patients as well. Obviously, the type of work and contact with and proximity to others are major concerns here, and need to be dealt with on a case-by-case basis. It is courteous to take precautions if patients work in an environment in which they are monitored (e.g. nuclear power plant workers, drivers of garbage trucks, visitors to certain federal buildings, travelers through certain airports) or in which the trash is monitored (e.g., hospitals which attempt to thwart sanitary landfill radiation detectors).

Urinary incontinence is a problem with most therapy radiopharmaceuticals, and needs to be ascertained and dealt with. Some patients may be depended upon to appropriately manage catheters at home, but others may not.

CALCULATIONS

The methodology for calculation of absorbed dose from an external source, such as a patient or a spill on the floor, was described in 1970 in Appendix I of NCRP 37. Other methods were suggested in the more recent Appendix U of NUREG-1556 Vol. 9 (the NRC guidance associated with the patient release rule pursuant to 10 CFR 35.75), NCRP Report No. 155, a Society of Nuclear Medicine (SNM) guidance document, and other references at the end of this document. The method of NCRP Report No. 37 uses a relatively simple approach, and gives reasonable estimates of dose, better than the simple release criterion based on retained activity that was used for many years, but whose origin is uncertain. In a 2007 article in Health Physics, we argued that licensees rely too heavily on the overly conservative approach in NUREG-1556, and should thoughtfully adopt more realistic methods in most cases. The recently issued NRC guidance addressing the post-radioiodine therapy release of patients who might come into contact with infants and/or young children was discussed in a Journal of Nuclear Medicine Newsline article in 2008, which concluded that there are no data to support such guidance.

The method can certainly accommodate increasing complexity. For example, current research with more realistic body phantoms provides dose factors that do not rely on treating patients as point sources, which is somewhat conservative in many cases,.  The specific gamma ray constant as used in most dose calculations assumes that the patient is an unshielded point source. Patients are of course not unshielded point sources, and if biodistribution and shielding can be measured or calculated, a more appropriately determined or modified gamma ray constant should be used.

When the external source is a fixed radioactive source, like a spill on the floor, the physical half-life of the involved radionuclide will obviously apply and should be used in any dose scenario. When the external source is a patient with a radiopharmaceutical, one should always combine the biological half-life with the physical half-life to get the effective half-life. The effective half-life should be used in the dose equation instead of the physical half-life when the radiation source is a patient.

Teffective=Tbiological×Tphysical  Tbiological +Tphysical

THE USEFUL DOSE EQUATION

The equation that can be used in most cases to estimate dose to others from a released patient is:

D (∞)=34.6×0.25×A× T12× Γ 1−e−0.693 ×∞T12  (100 cm)2D (∞)=8.66× 10−4A× T12× Γ

See Appendix I for the derivation of this expression. Appendix II shows a comparison of a dose calculation for a thyroid cancer patient using this expression and using the NUREG 1556 equation.

Since the T1/2 or Teff is representative of the clearance from the radioactive patient’s total body, this version of the equation has one component with respect to half-time, e.g., if there is a monoexponential total body clearance of radiopharmaceutical. If total body clearance follows more complex kinetics, e.g., biexponential or triexponential clearance, then additional components are necessary. In this case (T1/2)TB or (Teff)TB must be replaced by i fi (Ti)eff where fi is the uptake fraction of the i-th exponential component and (Ti)eff is the effective half-time of the i-th exponential component.

One can use a model with fixed parameters to facilitate dose calculation, but of course actual measurement of such parameters for individual patients will further improve the accuracy of the dose estimate. For example, in the case of radioiodine treatment of thyroid cancer and hyperthyroidism, a three-component model with a set of assumed uptake fractions and effective half-times has been advanced in NUREG-1556, Vol 9Error: Reference source not found, representing an eight-hour non-void period with an occupancy factor of 0.75, and non-thyroidal and thyroidal compartments. Thus, after the non-void period, the total body clearance is biexponential in this case. In Appendix II we show an example of this approach, but we also note that there are no scientific data to support this non-void componentError: Reference source not found.

The examples given in this document employ certain assumptions (e.g. in Example 5, “Assume that a person spends one fourth of all time at 1 meter from the subject (very conservative)”) that are needed to complete the calculations. As with model parameters (e.g. effective half-times), in the absence of any information, default values may be used, but it is always best to apply reasonable values for each individual case, based on information gleaned in discussions with each patient. Assigning an occupancy factor of 0.25 (one fourth of all time after patient release) at 1 m from the patient is probably conservative for most situations, and lower values may be applied if appropriate. Of course, one can also envision family situations in which this is quite reasonable, or in which even higher values of occupancy might reasonably be applied. Use good judgment in each case. When there is uncertainty, being conservative is always the preferred approach, but being unreasonably conservative may place burdens on patients and their families, as well as the treating institution.

Another component of dose considered by some is an internal dose component, from radioactivity released from the patient in saliva, sweat, or other body fluids and accidentally taken in by a person who may have also received an external radiation dose. In the 2007 HPJ paper, we evaluated this component and concluded that the internal dose component can be “considered to be negligible due to the use of an intake factor of 10-6”and that for individuals exposed to either thyroid cancer or hyperthyroid patients “the internal dose component does not have to be taken into account, as it will always be less than 10% of the external dose component.”Error: Reference source not found.

SPECIFIC GAMMA RAY CONSTANTS

FOR COMMON RADIONUCLIDES AS UNSHIELDED POINT SOURCES

18F			5.73 R-cm2/mCi-h

131I			2.2 R-cm2/mCi-h

99mTc		0.78 R-cm2/mCi-h

67Ga			0.79 R-cm2/mCi-h

201Tl			4.7 R-cm2/mCi-h

137Cs			3.3 R-cm2/mCi-h

60Co			13.2 R-cm2/mCi-h

111In			3.2 R-cm2/mCi-h

123I			1.6 R-cm2/mCi-h

89Sr			0.00046 R-cm2/mCi-h

192Ir			4.8 R-cm2/mCi-h

153Sm		0.45 R-cm2/mCi-h

In some cases, modified versions of these values may be used. For example, when 131I is dispersed in a patient, the value is on average approximately 0.6 times that of the unshielded specific gamma ray constant, (about 1.3 R-cm2/mCi-h). This was measured in 157 cases with administration of 131I-Bexxar. It should be noted that this factor was shown to be a function of patient mass. This factor will be applied to the extrathyroidal component of 131I-NaI subjects in the calculations below. For the thyroidal component, the activity is close to the surface of the body, and use of the unshielded point source value is appropriate.

PHYSICAL HALF-LIVES FOR COMMON RADIONUCLIDES (days)

18F				0.076

131I				8.04

99mTc			0.25

67Ga			3.26

201Tl			3.0

137Cs			11,000

60Co			1,920

111In			2.8

123I				0.55

89Sr			50.5

192Ir			73.8

153Sm			1.93

EXAMPLE

Here is an example dose calculation using for a thyroid cancer patient receiving a 100 mCi Na131I activity treatment prescription:

Occupancy Factor = 0.25

F1 	= 	0.95 (extrathyroidal fraction)

T1eff 	= 	0.32 d (effective half-time for extrathyroidal fraction)

F2 	= 	0.05 (thyroid fraction)

T2eff 	= 	7.3 d (thyroid effective half-time)

D(∞)=34.6 hd 2.2rem−cm2mCi−h (100 mCi)  (100 cm)2×0.25 0.6 ×0.95 ×0.32 d+0.05× 7.3 dD(∞)=8.66x10−4× 100 mCi×2.2rem−cm2mCi−h× 0.6 ×0.95 ×0.32 d+0.05× 7.3 dDDE(∞)=0.104 rem=104 mrem

EXAMPLE PROBLEMS

1.	A patient has been given 5 mCi 67Ga citrate for imaging infection. Three days later a surgeon, wishing to operate, hesitates to do so fearing radiation exposure. The surgical procedure would take about 3 hrs. Estimate the radiation dose to the surgeon, and compare it with common radiation exposure sources.

The T1/2 of 67Ga is 78 hrs. Very little is excreted over the 3-day period, so use of the physical half-life is reasonable. Let us assume 2.5 mCi is still in the patient. Assume that the average distance between the patient and the surgeon is 2/3 meter (this is conservative; 1 meter might also be considered)

D=34.6 (2.5 mCi) (3.26 d) 0.79rem−cm2mCi−h 1−e−0.693 324 d3.26 d  (67 cm)2D=0.0013 rem=1.3 mrem

For comparison, background = 1 mrem/day on average in USA, about 1.6 mrem/day in Denver. In airplanes, one receives about 1 mrem/1000 mi flown. Natural background averages 300 mrem/y average, + about 300 mrem/y currently from man-made radiation. So one could say that this exposure could have resulted from exposure to (1.3/300) x 365  2 days of natural background radiation.

2.	An ultrasound technologist balks at performing a gallbladder procedure in a patient who has just had a 99mTc-disofenin ("HIDA") scan. The patient received 2 mCi 99mTc 3 hrs. ago. What is the approximate radiation dose to the technologist assuming that the ultrasound examination takes 30 minutes? Again, assume a distance of about 2/3 of a meter:

A=2 mCi ×e−0.693 324 d0.25 d=1.4 mCi

Activity remaining:

D=34.6 (1.4 mCi) (0.25 d) 0.78rem−cm2mCi−h 1−e−0.693 0.524 d0.25 d  (67 cm)2D=0.0001 rem≈0.1 mrem

3.	A patient swallows 400 μCi 123NaI for a thyroid study and soon vomits a small amount on the floor. A calibrated ion chamber measures 0.2 mR/hr 50 cm above the vomitus. About how much activity did the patient vomit?

0.0002 rem=34.6 (Q0 mCi) (0.55 d) 1.66rem−cm2mCi−h 1−e−0.693 124 d0.55 d  (50 cm)2Q0=0.33 mCi=330 μCi, so the patient vomited most of the administered activity

4.	A technologist inadvertently spills 1 mCi of 99mTc on the floor while injecting a patient with 25 mCi of 99mTc-MIBI. The technologist does not clean up the spill. What is the most radiation that could reasonably be absorbed by the most exposed person frequenting that area? Let’s assume an average distance of 1 m for this calculation and an exposure time of 4 hours:

D=34.6 (1 mCi) (0.25 d) 0.78rem−cm2mCi−h 1−e−0.693 424 d0.25 d  (100 cm)2D=0.00025 rem=0.25 mrem

5.	A patient requires 50 mCi 131I for a Graves' disease therapy. Her thyroid uptake is 55%, her thyroid biological half-life is 5 days, and her renal function is normal. If you treated her as an outpatient, what is the highest expected radiation dose to someone with whom she shares a household? Assume she sleeps alone and does not share eating utensils. Assume that a person spends one fourth of all time at 1 meter from the subject (very conservative). Activity not taken up by the thyroid is assumed to have a 0.25 d biological half-life, which works out to be a 0.24 d effective half-time.

Teffective=5 d×8 d5 d+8 d=3.1 dDthyroidal=34.6 (0.55×50 mCi) (3.1 d) 2.2rem−cm2mCi−h 1−e−0.693 3.1 d×0.25  (100 cm)2Dnon−thyroidal=34.6 (0.45×50 mCi)(0.24 d) 1.3rem−cm2mCi−h 1−e−0.693 0.24 d×0.25  (100 cm)2D=Dthyroidal+Dnon−thyroidal=0.162+0.006 rem=0.17 rem=170 mrem

Note: We used the estimate of 1.3 R-cm2/mCi-h for the non-thyroidal component, as described on page 4.

Note: Since we integrated to infinity, the (1-e) terms become 1.0, and could have been left out.

Note: The biological half-life of the non-thyroidal component in a person with normal renal function is about 8 hrs or 0.33 days, and the effective half-life is 0.32 days.  In a hyperthyroid person with normal renal function it is about 6 hours or 0.25 days, and the effective half-life is 0.24 days. In a hypothyroid patient with normal renal function it is about 12 hours or 0.5 days, and the effective half-life is 0.47 days.

6.	A thyroid cancer patient requires 200 mCi 131I for therapy. Assume that uptake in the postoperative thyroid remnants is 1% and that the patient has normal renal function. If you treated him as an outpatient, what is the highest expected radiation dose to someone with whom he shares a household? The biological half-life of iodine in normal thyroid tissue is 80 days. If the patient is treated with Thyrogen (recombinant human TSH) assume a normal biological halflife of non-thyroidal iodine of 0.33 days.  If the patient went naturally hypothyroid, assume a biological half-life of non-thyroidal iodine of 0.5 days.  Here we assume rhTSH treatment.

Teffective=80 d×8 d80 d+8 d=7.3 dDthyroidal=34.6 (0.01×200 mCi) (7.3 d) 2.2rem−cm2mCi−h ×0.25  (100 cm)2Dnon−thyroidal=34.6 (0.99×200 mCi)(0.32 d) 1.3rem−cm2mCi−h ×0.25  (100 cm)2D=Dthyroidal+Dnon−thyroidal=0.0278+0.0712 rem=0.099 rem≈100 mrem

Note: we did not show the (1-e) terms this time, since they are equal to 1.0.

7.	A radiation inspector comes through your department and detects 10,000 dpm from a wastebasket in a public area. A patient injected for a 99mTc-MDP bone scan has been sitting near by drinking water as directed, and the activity appears to be coming from saliva on a foam cup he has discarded. Calculate the radiation dose to the most exposed person from this wastebasket, assuming the cup is left there. How would you respond to the radiation inspector? Assume that the person sits 1 m from the source for 4 hours (again very conservative). There is no regulation preventing Tc-99m patient-contaminated articles from being thrown in general trash.

A=10,000dismin2.22x1010disminmCi=4.5x10−6 mCiD=34.6 (4.5x10−6 mCi) (0.25 d) 0.78rem−cm2mCi−h 1−e−0.693 424 d0.25 d  (100 cm)2D=1.1x10−9 rem=1 nrem (nanorem)=negligible!

8.	A patient receives 100 mCi 153Sm-EDTMP and is driven home from the nuclear medicine department in a car or taxi. The trip takes 2 hours; the patient sits in the back of the vehicle on the right. Calculate the dose to the driver.

D=34.6 (100 mCi) (1.93 d) 0.45rem−cm2mCi−h 1−e−0.693 224 d1.93 d  (200 cm)2D=0.0022 rem=2.2 mrem

APPENDIX I: DERIVATION OF THE DOSE EQUATION

The exposure rate at a fixed distance r from a point source of radioactivity due to photon emissions is given as:

X=AΓr2

X

is the exposure rate, A is the source activity,  is the specific gamma constant for the radionuclide, and r is the distance from the source. Traditionally, the units of  have been R-cm2/mCi-hr, the SI unit equivalent is (C/kg)-m2/MBq-h. To use the traditional units, we would introduce A in mCi and  r in cm. To evaluate the exposure, X, over a fixed time period T when the source strength does not change with time, we would just multiply the exposure rate by the exposure time, T:

X=A Γ Tr2

D (0→t)=0TA e−0.693 tT12 Γ r2 dtD (0→t)=A e−0.693 tT12 Γ 0.693 T12 r2=1.443A T12 Γ 1−e−0.693 tT12  r2But, since the source activity in actuality is decreasing with time, due to simple radiological decay, biological removal or a combination of the two, A is not a constant, and we have to integrate over time, t, to get the total exposure X. Also, since 1 R  1 rad  1 rem for photons, we can equate exposure X with dose D, to obtain:

where D (0t) is the total accumulated dose from time 0 to time t.

THE USEFUL DOSE EQUATION

We can simplify the expression above so that a simpler final expression necessary for dose estimation is not that complicated.

- If we use half-lives in days instead of in hours, we multiply the previous equation by 24 hours/day.

- Also, we need to account for the fact that members of the public are not likely to spend 100% of the day in proximity to the patient. This can be done through use of an occupancy factor, OF; an occupancy factor of unity is generally assumed for short exposure periods whereas a factor of 0.25, signifying that it is unlikely that an individual will spend more that 25% of the day in proximity to a released patient, is generally applied and acceptable for longer periods of exposure.

- We can thus use the following expression for our dose calculations for patients:

D (0→t)=1.443×24×OF×A× T12× Γ 1−e−0.693 tT12  r2D (0→t)=34.6 ×OF×A× T12× Γ 1−e−0.693 tT12  r2

One should generally use an effective half-life (which accounts for biological clearance of the radionuclide as well as radioactive decay) for T1/2 instead of a physical half-life. This is easily calculated as:

Teffective=Tbiological×Tphysical  Tbiological +Tphysical

This expression thus gives the dose integrated to time t to a person at distance r from the patient. This is a very useful dose equation applicable to a variety of scenarios, as illustrated in the Examples section.

Further simplifications can be made if we:

- integrate the dose equation to infinite time, i.e., to get the potential total dose to others due to total decay of radioactivity in the patient), D(∞), and

- assume an occupancy factor of 0.25 and a standard distance of 1 m from the patient to the subject, which is a distance generally used for patient release calculations to ensure compliance with the 5 mSv dose limit (note that it may be appropriate to use other occupancy factors depending upon the specific circumstances):

D (∞)=34.6×0.25×A× Teff× Γ 1−e−0.693 ×∞Teff  (100 cm)2D (∞)=8.66× 10−4A× T12× Γ

If we have more than one component of biological removal (e.g. a short biological half-time and a long biological half-time, which is usually true for NaI, for example), we can associate the total activity to each component using fractions F1 and F2, associated with Teff-1 and Teff-1:

- D (∞)=34.6×0.25×Γ×A×F1 ×Teff−1 1−e−0.693 ×∞Teff−1+F2 ×Teff−2 1−e−0.693 ×∞Teff−2  (100 cm)2D (∞)=8.66× 10−4A× Γ×F1 ×Teff−1+F2 ×Teff−1

APPENDIX II: COMPARISON OF NRC AND HPJ SOLUTIONS TO EXAMPLE THYROID CANCER PATIENT RELEASE

Here is an example dose calculation using two methods, and variables as defined in NUREG-1556, Vol. 9Error: Reference source not found and our 2007 HPJ article for a thyroid cancer patient receiving a 100 mCi Na131I activity treatment prescription. A three-component model with a set of assumed uptake fractions and effective half-times was used in NUREG-1556, Vol 9Error: Reference source not found, assuming an eight-hour non-void period with an occupancy factor of 0.75, and then two exponential terms representing non-thyroidal and thyroidal compartments. In our HPJ article, we showed that one can omit the non-void period, as there are no scientific data to support such a dose component, and modify the specific gamma ray constant by a factor of 0.6 to account for attenuation and activity distribution of the Na131I in the body (i.e., the extrathyroidal component, see page 4 above).

E1 	= 	0.75 (NUREG assumed occupancy during non-void period)

TNV 	= 	0.33 d (non-void period)

E2	= 	0.25 (NUREG assumed occupancy after non-void period)

F1 	= 	0.95

Teff-1 	= 	0.32 d

F2 	= 	0.05

Teff-2 	= 	7.3 d

D(∞)=34.6hd 2.2rem−cm2mCi−h (100 mCi)  (100 cm)2×0.75 ∗8.04 d ∗(0.8) [1−e−0.693(0.33 d8.04 d)]+0.25×0.95× 0.32 d×e−0.693(0.33 d8.04 d) +0.25×0.05× 7.3 d×e−0.693(0.33 d8.04 d) DDE(∞)=0.227 rem=227 mremNRC equation:

D(∞)=34.6 hd 2.2rem−cm2mCi−h (100 mCi)  (100 cm)2×0.25 0.6 ×0.95 ×0.32 d+0.05× 7.3 dD(∞)=8.66x10−4× 100 mCi×2.2rem−cm2mCi−h× 0.6 ×0.95 ×0.32 d+0.05× 7.3 dDDE(∞)=0.104 rem=104 mremHPJ equation:

As can be seen the more realistic calculation provides a dose estimate that is a factor of two lower than the NRC guidance approach, which is overly conservative.

REFERENCES

Marcus CS and Aldrich R. Avoiding Solid Waste Contamination Problems from Iodine-131 Patients. J.Nucl.Med.38(9):26N,1997

Siegel JA, Sparks RB. Radioactivity appearing at landfills in household trash of nuclear medicine patients: much ado about nothing? Health Phys. 2002; 82:367-372.

National Council on Radiation Protection and Measurements. Precautions in the Management of Patients Who Have Received Therapeutic Amounts of Radionuclides, NCRP Report 37, October, 1970.

U.S. Nuclear Regulatory Commission. Specific Guidance About Medical Use Licenses (NUREG-1556, Vol. 9, Rev. 2) Final Report. 2008.

U.S. Nuclear Regulatory Commission. Code of Federal Regulations Part 35, Subpart C Release of individuals containing unsealed byproduct material or implants containing byproduct material. 10CFR35.75.

National Council on Radiation Protection and Measurements. Management of Radionuclide Therapy Patients. NCRP Report No. 155, 2006.

Siegel JA. Nuclear Regulatory Commission Regulation of Nuclear Medicine: Guide for Diagnostic Nuclear Medicine and Radiopharmaceutical Therapy. Reston, VA: Society of Nuclear Medicine; 2004.

Siegel JA. Tracking the origin of the NRC 30-mCi rule. J Nucl Med 2000; 41:10N-16N.

Siegel JA, Marcus CS, Stabin MG. Licensee over-reliance on conservatisms in NRC guidance regarding the release of patients treated with 131I. Health Phys. 2007; 93(6):677-677.

Siegel JA, Silberstein EB. A closer look at the latest NRC patient release guidance. J Nucl Med. 2008; 49:17N-20N.

Siegel JA, Marcus CS, Sparks RB. Calculating the absorbed dose to others from the radioactive patient: Line source model versus point source model. J Nucl Med. 2002; 43:1241-1244.

Sparks RB, Siegel JA, Wahl RL. The need for better methods to determine release criteria for patients administered radioactive material. Health Phys. 1998; 75:385-388.

Siegel JA, Kroll S, Regan D, Kaminski MS, Wahl RL. A practical methodology for patient release after tositumomab and 131I-tositumomab therapy. J Nucl Med. 2002; 43:354-363.

Cameron, J.R., A radiation unit for the public. Physics and Society 20 (1991) 2.